## Slide 1: Introduction
By: Guy Soffer

## Slide 2
LEGACY STATION CONFIGURATION

## Slide 3
NEW STATION CONFIGURATION
DMM
Agilent 34401A
PRT/SPRT
UNIT
Any computer with NI-GPIB-USB-HS
Replaces the legacy 3714
1x
GPIB cable
Optional
By: Guy Soffer

## Slide 4: Thermal-Calibrator Application Overview

## Slide 5
A
B
C
D
E
Report viewer
Procedure viewer
Station configuration
Main screen
Manual control
By: Guy Soffer

## Slide 6
Menus to access
files and tools
Station configuration
Real time status and control
Real time messages for debug
By: Guy Soffer

## Slide 7
By: Guy Soffer

## Slide 8: STATION CONFIGURATION

## Slide 9
DMM
PRT/SPRT
UNIT
Station configuration file
Device detection tool
Note:
Station configuration editor / display
By: Guy Soffer

## Slide 10
By: Guy Soffer

## Slide 11
UNIT data structure
UNIT configuration source file (JSON)
UNIT information
GPIB bus#
0 if a single GPIB-USB-HS is used
Device#
KNC units are 1 by default
Red background indicates that the date expired
By: Guy Soffer

## Slide 12
DMM data structure
DMM configuration source file (JSON)
DMM information
Should match the DMM configuration
Should match the DMM configuration
“Fluke” or “SCPI”
By: Guy Soffer

## Slide 13
PRT data structure
PRT configuration source file (JSON)
PRT information
Should match the SPRT calibration report
By: Guy Soffer

## Slide 14: STATION OPERATION

## Slide 15
If connection successful
Operation buttons
disabled
Click
Connect
Click disconnect
Operation buttons
enabled
By: Guy Soffer
DMM and unit were detected

## Slide 16
Connect / disconnect
Make the unit BEEP
Enable / disable the unit’s printer
Abort running procedure and turn off the unit’s well
Procedures
Unit’s temperature minus SPRT
Units selector
Skip to the next temperature point
Show the report
Show the ramp-and-soak program
Unit’s status
Temperatures
By: Guy Soffer

## Slide 17: MONITOR or MANUAL CONTROL

## Slide 18
The “MONITOR” mode supports stations with unit only, SPRT only, or both connected configurations.
The TARGET and UNIT fields will be updated if the unit is present on the GPIB bus.
The SPRT/UUT field will be updated if the DMM is present on the GPIB bus.
If the unit is present on the GPIB bus, it can be manually controlled via the “Manual control” dialog box (under the “Config” menu.
By: Guy Soffer

## Slide 19: STANDARDIZATION

## Slide 20
Both the unit and DMM must be present on the GPIB bus.
The SPRT/UUT field will be updated if the DMM is present on the GPIB bus.
Note: The unit must be place an CAL mode with activated IEEE mode. Refer to unit’s user manual to further instructions.
Advance to the next standardization point.
Show the standardization report.
Not applicable in standardization mode.
By: Guy Soffer

## Slide 21: CALIBRATION

## Slide 22
Both the unit and DMM must be present on the GPIB bus.
The SPRT/UUT field will be updated if the DMM is present on the GPIB bus.
Advance to the next calibration point.
Show the calibration report.
Show the calibration procedure steps.
The calibration protocol (target points, dwell times, and pass/fail criteria) is in a standard “Ramp-and-Soak” program <3605A_fullCalib_F.json>.
A quick calibration procedure can be run by loading the program <3605A_quickCalib.json> before starting the procedure. Different calibration procedures can added as well.
By: Guy Soffer

## Slide 23: AUTOMATIC STANDARDIZATION AND CALIBRATION

## Slide 24
Both the unit and DMM must be present on the GPIB bus.
First the standardization procedure will run.
Than, the calibration procedure will begin.
The natural cooling time of 5 min is inserted between the end of the STD to the beginning of the CAL.
Advance to the next standardization / calibration point.
Show the current procedure report.
Show the calibration procedure steps.
By: Guy Soffer

## Slide 25: RAMP AND SOAK PROGRAM

## Slide 26: User defined test protocols (Ramp-and-Soak)
Load a program and start a “Ramp And Soak” process
Example of a quick calibration test program
<3605A_quickCalib.json>
Automatically transition between test points
Soak time after unit’s temperature is stable
How many times to repeat the sequence
The test sequence
Pass / fail criteria (°C)
Soak / dwell time (minutes)
Target temperature (°C)
Not implemented
(°C/minute)
Note: All temperature values must be entered in °C units.
By: Guy Soffer

## Slide 27
Depend on Ramp-and-Soak program configuration, only the unit or both the unit and DMM must be present on the GPIB bus.
Advance to the next target point.
Show the test report.
Show the test procedure steps.
By: Guy Soffer

## Slide 28: WELL PERFORMANCE TESTING

## Slide 29
UNIT
Any computer with NI-GPIB-USB-HS
WELL TESTING
STATION CONFIGURATION
Note: DMM and SPRT are optional and no being used during the well testing procedure.
By: Guy Soffer

## Slide 30: Protocol for testing the thermal performance
Chart: Temperature profile
| 0 | 31 | 99 | 191 | 283 | 405 | 527 | 709 | 891 | 1134 | 1207 | 1299 | 1391 | 1483 | 1575 | 1697 | 1819 | 2062 | 2123 | 2305 | 2547 | 2879 | 3481
 | 77.153 | 84.911 | 102.326 | 119.7644 | 136.1966 | 153.8978 | 168.8198 | 186.8234 | 200.9858 | 212.0432 | 193.8254 | 175.7372 | 158.6462 | 141.3086 | 124.4264 | 105.4454 | 88.1456 | 76.9658 | 68.9828 | 51.2114 | 32.5256 | 15.1214 | 139.8722
 | 76.82 | 83.84 | 98.6 | 114.44 | 128.48 | 146.12 | 161.06 | 178.88 | 195.8 | 211.82 | 191.84 | 173.3 | 154.76 | 135.68 | 116.6 | 98.24 | 78.98 | 76.82 | 66.56 | 47.84 | 29.12 | 10.76 | 135.68
 | 77.18 | 88.16 | 109.4 | 129.56 | 151.52 | 169.88 | 190.94 | 209.12 | 228.2 | 212.18 | 196.16 | 178.7 | 161.24 | 144.32 | 127.4 | 109.76 | 93.02 | 77.18 | 69.44 | 52.16 | 34.88 | 17.24 | 144.32
Load a program and start a “Ramp And Soak” process
Results of good well
Red lines indicate minimum and maximum acceptance criteria
Black line indicates the measured well’s temperature
Results of defective well
Chart: Temperature profile
| 0 | 32 | 100 | 192 | 284 | 406 | 528 | 710 | 892 | 1134 | 1207 | 1299 | 1391 | 1483 | 1576 | 1697 | 1819 | 2062 | 2124 | 2306 | 2547 | 2880 | 3482
 | 77.1494 | 84.8516 | 102.164 | 118.8662 | 132.3878 | 147.7994 | 160.7792 | 176.4266 | 188.5622 | 200.3288 | 193.2062 | 175.7822 | 158.1782 | 141.359 | 124.43 | 106.4048 | 89.8952 | 76.946 | 69.3194 | 53.8754 | 40.631 | 28.9346 | 139.6742
 | 76.82 | 83.84 | 98.6 | 114.44 | 128.48 | 146.12 | 161.06 | 178.88 | 195.8 | 211.82 | 191.84 | 173.3 | 154.76 | 135.68 | 116.6 | 98.24 | 78.98 | 76.82 | 66.56 | 47.84 | 29.12 | 10.76 | 135.68
 | 77.18 | 88.16 | 109.4 | 129.56 | 151.52 | 169.88 | 190.94 | 209.12 | 228.2 | 212.18 | 196.16 | 178.7 | 161.24 | 144.32 | 127.4 | 109.76 | 93.02 | 77.18 | 69.44 | 52.16 | 34.88 | 17.24 | 144.32
<3605A_wellTest_NO_Insolation.json>
FREE = 1: Dwell timer starts immediately after setpoint is set
Deviation is calculated between unit and set-point
By: Guy Soffer

## Slide 31: Test Report

## Slide 32
Type of test
STD, CAL, MEAS
Station configuration during the test
Measurements at each temperature point
Average ambient temperature during the test
In case of standardization, the new calculated coefficients
Refresh the report of a running test
Read a saved report
Save the report to a JSON file
Generate an excel report file
Double click “Root” to expand
Note: The reports can displayed in °C or °F units. Select the desired units in the application’s main screen before opening the report viewer windows.
By: Guy Soffer

## Slide 33
Calibration report in excel format
Standardization report in excel format
Station
setup
Results
The new coefficients
Results
Station
setup
Note: The reports can be generated in °C or °F units, depends on the Report viewer window.

## Slide 34: INSTALATION

## Slide 35
1. Download and install the NI IEEE488.2 drivers for windows
https://www.ni.com/en-us/support/downloads/drivers/download.ni-488-2.html#467646
2. Download the ThermoCalibrator application
https://drive.google.com/file/d/1KcgN22rnByD-dbM9KaY9K4KyVqcFsqsT/view?usp=share_link
3. Unzip the ThermoCalibrator files into the desired folder

## Slide 36
4. Make a shortcut on the desktop. The link can include the station configuration file to load on start.
"D:\KNC\ThermalCalibrator\ThermalCalibrator_Station.exe" -file ./config/station/station_04_3605A.json
The path and file name of the station configuration file to load on start
Right click and select
“Properties”